Lakeshore CRX 6.5 K Cryoprobe Station

User Guide

Last updated: Feb. 16, 2026

Mailing list:

Allen152-lakeshore-cryo@lists.stanford.edu

For problems or training, please contact superusers:

Koustav Jana (kj2011@stanford.edu)

Tara Pena (tpena406@stanford.edu)

Zhengliang Bian (bianzl@stanford.edu)

Venting the chamber – Sample Load/Unload

- Confirm the probe arms are retracted from the sample stage.

- Close the pumptrain isolation valve by turning clockwise.

- Turn off the turbo pumping system using the software GUI. Do NOT touch the pump.

  - If you used the D-sub connector to disable ASDF, unplug it now. The D-sub connector shorts the scroll pump so that it will not turn off when the GUI command for stopping the turbo is sent.

  - Place it on top of the pump so that it does not go missing.

- Open outer lid by turning on N2 and unscrewing the screws. This lid should be set on a cleanroom wipe in its proper orientation (do not put the inside surface facing up as anything scratching it could cause bad vacuum).

- Turn on the N2 line counter-clockwise. Wait a minute or so until the chamber is at atmosphere. Keep N2 flowing while lid is open.

- Unscrew the inner lid. These screws will not run away but be careful.

- Load sample. Check that sample will be in field of view of microscope and that focus is in range.

- Move in probe arms that you will use. Move them close to where you want to probe, but give them space so they can move during pump down and cool down process.

- Close the inner lid. Screw in all screws finger tight (do not overtighten).

- Turn off N2 by closing clockwise.

- Close outer lid.

- Check the speed of the turbo in the software GUI. If the turbo speed Is still >35000 rpm, wait until it drops below 35000 rpm.

- If the turbo speed is <25000 rpm but not 0, VERY QUICKLY open (CCW) and close (CW) the manual vent valve. This will slow the turbo down faster, within a couple minutes.

- When the turbo speed is at 0 rpm, open manual vent valve on turbo.

- Once hissing ceases, close the turbo vent valve completely.

Pumping the chamber

- Replace and screw in all lids.

- Make sure N2 purge valve is closed, turbo is at 0 rpm, and turbo vent valve is closed.

- Make sure the turbo speed is 0 rpm in the software GUI.

- Open the pumptrain isolation valve by turning CCW.

- Open the turbo vent valve ¼ turn CCW.

- Open the N2 valve to purge the system. Purging with N2 allows faster pumping.

- Close the turbo vent valve.

- Close the N2 valve.

- Plug in the D-sub connector into the J6 connection to the left of the RS232 cable on the back of the pump. The D-sub connector shorts the scroll pump power supply so that it will maintain base pressure during measurement.

- Turn on the pump with the GUI. Monitor the pressure on the GUI. It should reach <1e-3 Torr within a few minutes.